# Replacing the Stirling cooler

Steps to remove the cooler

- Open the vacuum chamber-see protocol on this.

- Loosen the vacuum fastener/clump, remove the screws that connect the stirling cooler to the base holder. This way you can be able to gently move the copper base plate

Vacuum fastener/clump

Oasis cooler pipes

Stirling cooler

- Gently put aside the temperature sensor by unscrewing the screw on the side.

Temperature sensor screw

- Remove the pipes coming from the chillers (see image in 2). Be careful at this point because there might be cooling liquid. Have spare tissues or stoppers for this step.

- Remove the screws from the steel holder from the bottom of the TREX rack.

Screws at the bottom of the TREX

- You will then see cables from the Stirling cooler going into the sunpower blue box. For the two cables-orange is +OUT and white is -OUT. Cut the cables to have more space to move the stirling cooler. If you do have a better tool to access this, then you do not need to cut the cables.

Unscrew on these holes

Orange and white cables

- You then unscrew this from the holes on top of the box and remove the cooler completely.

Replacing the cooler:

- Have someone hold the Stirling cooler while you screw back the two cables to the SunPower blue box.

- Put the Stirling cooler back in holding place. Screw back the steel holder.

- Connect back the pipes from the chiller and fasten.

- Place back the temperature sensor back, then the copper base plate, while fastening the vacuum clamp and screw at the back to ensure alignment.

- See if it aligns with the linear actuator by adjusting the screw on the L corner (with the green, blue and red cables). Be careful with temperature sensor and any other cables.

- Close the vacuum chamber while ensuring that the screws align properly. You may need to adjust the L corner several times if you feel that the screws are a bit too tight or it is really hard to screw them in.